API Documentation

# Commands

A NULL value should be replaced with a 0 (char. Zero) when writing the command

## Prime

| OP | arg1 | var1 | arg2 | var2 | arg3 | var3 |
| --- | --- | --- | --- | --- | --- | --- |
| PRIME | CHEM_WASH | 1-8 (0,10) | -T | <sec> | -R | 36 |
|  | -C | 1-8 (10) |  |  |  |  |

### ARG1

| CHEM_WASH | The “CHEM_WASH” flag tells the system to prime wash to the corresponding channel (1-8) in var1 versus the reagent. If var1 is set to 0, this will prime the wash line but not each channels individual wash lines. If var1 is set to 10, this will prime all of the channels’ wash lines starting with 8 down to 1. |
| --- | --- |
| -C | The “-C” is a channel selection for reagents. Var1 will correspond to the reagents in channel 1-8. If var1 is set to 10, the system will prime all reagents. |

### ARG2

| -T | The “-T” sets the mode to be time based. This makes var2 an input in seconds. |
| --- | --- |

### ARG3

| -R | The “-R” sets the mode to be rate based. This should always be set to 36, any values outside of this are not tested. |
| --- | --- |

## Wash

| OP | arg1 | var1 | arg2 | var2 | arg3 | var3 |
| --- | --- | --- | --- | --- | --- | --- |
| WASH | -C | 1-8 | -T | <sec> | -R | 36 |
|  | CHEM_WASH | NULL |  |  |  |  |
|  | COLLECT |  |  |  |  |  |
|  | COMMON |  |  |  |  |  |
|  | ALL |  |  |  |  |  |

### ARG1

| CHEM_WASH | Washes all chemical channels. |
| --- | --- |
| -C | Washes a specific chemical channel based on VAR1 |
| COLLECT | Washes the common line all the way out to the collection/waste. |
| COMMON | Washes the common line up to the GFET chip. |
| ALL | Washes all runs (in order): CHEM_WASH, COMMON, COLLECTION |

### ARG2

| -T | The “-T” sets the mode to be time based. This makes var2 an input in seconds. |
| --- | --- |

### ARG3

| -R | The “-R” sets the mode to be rate based. This should always be set to 36, any values outside of this are not tested. |
| --- | --- |

## Purge

| OP | arg1 | var1 | arg2 | var2 | arg3 | var3 |
| --- | --- | --- | --- | --- | --- | --- |
| PURGE | ALL | NULL | -T | <sec> | -R | 36 |
|  | COMMON |  |  |  |  |  |

### ARG1

| ALL | Air purges from the air vent out to the collection/waste. |
| --- | --- |
| COMMON | Air purges from the air vent to before the GFET chip. |

### ARG2

| -T | The “-T” sets the mode to be time based. This makes var2 an input in seconds. |
| --- | --- |

### ARG3

| -R | The “-R” sets the mode to be rate based. This should always be set to 36, any values outside of this are not tested. |
| --- | --- |

## Collect

| OP | arg1 | var1 | arg2 | var2 | arg3 | var3 |
| --- | --- | --- | --- | --- | --- | --- |
| PRIME | -C | 1-8 (0,9) | -T | <sec> | -R | 36 |

### ARG1

| -C | Selects what channel to use for collection. 1-8: Normal reagent channels 0: Wash 9: Air |
| --- | --- |

### ARG2

| -T | The “-T” sets the mode to be time based. This makes var2 an input in seconds. |
| --- | --- |

### ARG3

| -R | The “-R” sets the mode to be rate based. This should always be set to 36, any values outside of this are not tested. |
| --- | --- |

## Pump

| OP | arg1 | var1 | arg2 | var2 | arg3 | var3 |
| --- | --- | --- | --- | --- | --- | --- |
| PRIME | -C | 1-8 | -T | <sec> | -R | 36 |

### ARG1

| -C | Selects which channel to pump from. |
| --- | --- |

### ARG2

| -T | The “-T” sets the mode to be time based. This makes var2 an input in seconds. |
| --- | --- |

### ARG3

| -R | The “-R” sets the mode to be rate based. This should always be set to 36, any values outside of this are not tested. |
| --- | --- |

## Wait

| OP | arg1 | var1 | arg2 | var2 | arg3 | var3 |
| --- | --- | --- | --- | --- | --- | --- |
| WAIT | -T | <sec> | NULL |  |  |  |

### ARG1

| -T | Var1 sets how long to wait for in seconds. |
| --- | --- |

# Example Script

%% Setup device PURGE ALL 0 -T 30 -R 36 PRIME CHEM_WASH 0 -T 10 -R 36 PRIME CHEM_WASH 8 -T 5 -R 36 PRIME CHEM_WASH 10 -T 10 -R 36 PRIME -C 10 -T 5 -R 36 WASH COMMON 0 -T 10 -R 36 WASH COLLECT 0 -T 10 -R 36 %% Start of sequence PUMP -C 1 -T 15 -R 36 WASH CHEM_WASH 0 -T 5 -R 36 WASH COMMON 0 -T 10 -R 36 PUMP -C 8 -T 15 -R 36 WASH CHEM_WASH 0 -T 5 -R 36 WASH COMMON 0 -T 10 -R 36 PUMP -C 5 -T 30 -R 36 WASH CHEM_WASH 0 -T 5 -R 36 WASH COMMON 0 -T 10 -R 36 PUMP -C 2 -T 30 -R 36 WASH CHEM_WASH 0 -T 5 -R 36 WASH COMMON 0 -T 10 -R 36 COLLECT -C 0 -T 100 -R 36

Text in black is a command. Text in red is a note.

# Script Writing Tips

- Washing after pumping chemicals

  - To prevent the chance of backflow, it is recommended to wash the chem channel that just pumped and then run a common wash after every pump operation.

- Proper first priming of wash

  - To properly prime the wash for all chemical channels, you should first PRIME CHEM_WASH using the argument 0 then 8 and finally you can run the full prime using 10.

- Proper first priming of chemical channels

  - After you run PRIME -C 10, you should run a full wash of the system to ensure there is no backwash. You should WASH CHEM_WASH and then WASH COMMON.

# Calibration Tips

- To figure out the correct timing for priming, it is recommended to test with de-ionized water.

- Prime times

  - Reset the System with an Air Purge:

    - Run the first calibration script to reset the system by performing an air purge for at least 100 seconds.

    - If liquid still appears in the output to waste after the purge, repeat the script until the system is fully cleared of liquid.

  - Prime the Test Chemical Channel:

    - Run the second calibration script to prime your test chemical channel for 100 seconds.

    - While the script is running, start a stopwatch.

    - Stop the stopwatch as soon as liquid starts flowing through the waste line.

    - The total time recorded on the stopwatch, plus an additional 2 seconds, will be the correct prime time for this chemical channel.

  - Repeat for Wash Prime Channel 8:

    - For the final step, prime Channel 8 with wash.

    - Follow the same procedure: reset the system with an air purge, then prime Channel 8 with wash for 100 seconds.

    - Start a stopwatch and stop it once liquid starts flowing through the waste line.

    - The total time recorded on the stopwatch, plus 2 seconds, will be the correct prime time for wash in Channel 8.

- Choosing your -R value

  - While the system was tested with -R 36, the usable range is anything below this. If the system is pumping fluid too fast, you can cut the speed roughly in half by dropping 36 to 18.

  - The -R value should not go above 36, this is the highest tested flow rate and anything higher might result in a blowout due to pressure build up in the fluidic system.